## Slide 1: Digital control and processing of optical frequency combs for precision measurements
Jean-Daniel Deschênes, Hugo Bergeron;
Université Laval, Canada
Laura Sinclair, Ian R. Coddington, Nathan R. Newbury
NIST Boulder, USA
CLEO 2017, San Jose, May 18th, 2017

## Slide 2: Talk outline
Digital control for combs
  Why digital?
  Available versions
  Implementation
  Features
Cycle slips in digital PLLs
  Cycle slip probabilities
  Effect of filter

## Slide 3: What do we mean by “digital PLL”
Replace as much electronics as possible by digital programmable hardware
DPLL platform consists of:
  Hardware platform with ADCs, DACs, FPGA
  DPLL Firmware running on FPGA
  PC software with a GUI
All software is open-source & free
Hardware is a mix of COTS and open “schematics”
Analog PLL
Digital PLL for laser locking:
Demod
Loop filter
ADC
Laser
DAC
Laser
Digital processing

## Slide 4: Why go digital?
Advantages
  Flexibility
  High integration possible
  Phase extraction in difficult conditions
  Replication is easier
Disadvantages
  Latency can limit BW
  Complexity
    Mostly at design/conception, not replication!

## Slide 5: Limitation of conventional analog phase detection
Microwave mixer used as phase detector
sin(Df) response: error Df must be < +/- p/2
y = sin(Df)
D1
Reference laser
Comb
LPF
Feedback
y
Df

## Slide 6: Why digital control for combs?
Combs are clockwork for optical clocks
Combs need to be locked to their optical clock
Particularities:
  low SNR
  High carrier frequency
    wide excursions possible
  Locking only a difference frequency
    lower requirements on lock and reference
  For a clock:
    No cycle slip!

## Slide 7
Optics package
Fb
Comb
PPLN
HNLF
D
Fceo
PLL
D
PLL
Reference laser
Fully self-referenced PM comb with digital locks
Ref: Sinclair, L. C. et al. "Invited Article: A compact optically coherent fiber frequency comb." Review of Scientific Instruments 86.8 (2015): 081301.
0.7 Litre

## Slide 8: Fully self-referenced PM comb with digital locks
Optics package
Fb
Peigne
PPLN
HNLF
D
Fceo
PLL
D
PLL
Reference laser
0.7 Litre
2.3 Litres
~200 Litres

## Slide 9
Optics package
Fb
Peigne
PPLN
HNLF
D
Fceo
PLL
D
PLL
Reference laser
0.7 Litre
2.3 Litres
Fully self-referenced PM comb with digital locks

## Slide 10: Features
2 Digital PLL:
  IQ phase demodulation + unwrapping: tolerant to perturbations
  PID loop filter
Built-in complete diagnostics:
  On-line phase noise PSD
  VNA
  Zero dead-time frequency counters
  Dither generator + lock-in
Clocked off comb itself: truly self-referenced
Factor ~100 reduction in volume

## Slide 11: Two hardware platforms
  COTS FPGA board
  +Open-source ADC/DAC board
    >=2.5k$
  USB2 interface
  2x 14-bits ADCs @ 100 MS/s
  3x DACs, 1 HV (55 V) output
    2x 16-bits @ 100 MS/S
    1x 24-bits @ 1 MS/s, HV (55V)
  COTS ADC/DAC/FPGA
    Cheap (~300$)
  Ethernet interface
  Runs Linux
  2x 14-bits ADCs @ 125 MS/s
  2x 14-bits DACs @ 125 MS/s
NIST Digital Lockbox
Red Pitaya
New!

## Slide 12: ADC noise & DAC noise
Quantization noise only: SNR = N*6 dB/bit + 1.76 dB
Time domain SNR = 10log10(Signal Power/Noise power)
  This noise is spread over the full Nyquist BW!
  Spectral domain SNR = Time domain SNR + 10log10(Fs/2/BW)
Example for Red Pitaya (14-bits ADC)
Still very high! ADCs are really good…
 | Time-domain SNR [dB]
Ideal 14-bits ADC | 86 dB
Red Pitaya’s ADC chip | 73 dB
Red Pitaya system | 63 dB
In 100 kHz @ 100 MS/s:
30 more dBs...

## Slide 13: High-level block diagram
2 independent DPLLs
+ Extensive diagnostics & logging to PC...
Demod
Loop filter
Freq counter
ADC
DAC
VNA
Lockin
PLL Channel 2
PLL Channel 1
DAC
ADC
(...)
FPGA

## Slide 14: Phase/frequency demodulator
Phase unwrapping
Selectable low-pass filters (7 MHz, 30 MHz, 60 MHz)
Tunable reference frequency (0-60 MHz)

## Slide 15: Phase/frequency demodulator: Phase extraction process
  Frequency shift by -Fref
  Filter
  Arctan(Q/I)
  Unwrap
Spectrum
f0-fref
Re{}=I
Im{}=Q
+f0
-f0

## Slide 16: Phase/frequency demodulator: Phase unwrapping
  Frequency shift by -Fref
  Filter
  Arctan(Q/I)
  Unwrap
Re{}=I
Im{}=Q
+p
-p
-p
+p

## Slide 17: Loop filter
Proportionnal, Integral and double-integral branch
  High noise rejection at low freq
Gain tuning through Bode diagram in UI

## Slide 18: Feature: Vector Network Analyzer
Measures baseband OR FM transfer functions
Very useful for loop tuning

## Slide 19: Frequency counter & lock - Example results – 5 days locked
CEO lock on a fiber PM comb
In-loop count
ILX Amplifier pump driver: E511!

## Slide 20: Cycle slips: What’s a cycle slip?
In a traditional counter:
In the complex plane:
Extra cycle!
Missing cycle!
Re{}=I
Im{}=Q
Normal counting
Normal counting,
signal’s phase is 0
w/ noise
Re{}=I
Im{}=Q
Cycle slip!
Re{}=I
Im{}=Q

## Slide 21: Cycle slips in complex plane
A cycle slip/unwrap error occurs when:
  Noise vector > signal amplitude
Prob[Cycle slip] = Prob[Noise > Signal] = erf(SNR)
Cycle slip!
Re{}=I
Im{}=Q
Black: Signal
Red: Noise

## Slide 22: Cycle slips in DPLL
SNR is a critical parameter for rate of cycle slips
Ref: Sinclair, L. C. et al. "Invited Article: A compact optically coherent fiber frequency comb." Review of Scientific Instruments 86.8 (2015): 081301.

## Slide 23: Optimizing SNR in DPLL
Want highest SNR for lowest prob cycle slip
How to choose filter BW in a DPLL?
SNR increases by decreasing BW
  Converges to a limit value when BW < Linewidth
Assumptions:
10 kHz Lorenzian beat
Fixed noise level

## Slide 24: SNR limit at low BW
SNR converges to a maximum value
  When BW < Linewidth
  Reducing BW further removes as much signal as noise
Linewidth = 10 kHz
Decreasing BW
Noise power
Signal power

## Slide 25
SNR=15dB
  Right on the edge

## Slide 26
SNR=24dB
  Ideal case

## Slide 27
SNR=33dB
  Barely any signal lost
  But too much AM noise!

## Slide 28
SNR=36dB
  Lost only 2 dB of signal on average
  But too much AM noise!
  IQ eye completely closed

## Slide 29: Probability of a cycle slip vs filter bandwidth
Can barely tolerate any signal loss at all before slip probability becomes catastrophic
SNR does not capture this effect
Desired operation zone
Linewidth = 10 kHz

## Slide 30: Conclusion
Presented an open-source digital PLL for combs locking
  Works on open-source hardware
  Now also available on an affordable commercial board
  You can run it in your lab for < 500$
Cycle slips
  High SNR is required for low cycle slips
  filter BW >> Linewidth (do not maximize SNR)

## Slide 31: Thank you!
Funding:
Links
  NIST digital lock box
    https://www.nist.gov/services-resources/software/fpga-based-digital-control-box-phase-stabilization-frequency-comb
  https://github.com/nist-ionstorage/digital-servo
  jean-daniel.deschenes@gel.ulaval.ca

## Slide 32: Extra slides

## Slide 33: ADC noise
Largest signal:
  2^N codes peak-peak,
  2^N/(2*sqrt(2)) codes RMS
Quantization noise:
  1/sqrt(12) codes RMS
SNR = 20*log10(2^N/(2*sqrt(2))*sqrt(12)) = N * 6 dB/bit + 1.76 dB

## Slide 34: DAC noise
Ideal quantization noise is level as ADC (0.3 codes RMS), BUT
Noise appears after the loop filter: in-loop!
  It gets rejected inside the lock BW as much as the laser noise
Demod
Loop filter
ADC
Quantizing noise
Laser
Freq noise
Hlaser
DAC model
Laser model
If quantizing noise is << laser noise at any point, then it’s negligible
Quantizing noise * Hlaser can be directly compared to the laser frequency noise

## Slide 35: DAC noise: Red Pitaya
In practice, DACs are usually closer to ‘ideal’ behavior
  Red Pitaya’s output is within 11.5 dB of ideal 14-bits DAC
Experience shows that 14 bits is plenty…
Hlaser
DAC Quantizing noise
DAC Quantizing noise

## Slide 36: Red Pitaya’s on-board oscillator
125 MHz Quartz oscillator
Stability:
  +/- 50 ppm stability (6 kHz) over Temp & aging
Phase jitter:
  1 ps (12 kHz-20 MHz)
Can be bypassed through modification (untested)

## Slide 37: Actual stability is much better: < 2 Hz over 12 hours
In fairly stable laboratory conditions
Red Pitaya internal clock vs double-oven OCXO
<2 Hz

## Slide 38: Actual measured phase noise
-70 dBc/Hz at f=1Hz
-125 dBc/Hz floor due to measurement limit
Phase noise PSD

## Slide 39: Feature: Dither lock-in
Gives continuous knowledge of loop sign & gain
Loop sign gives:
  Guarantees stable lock
  Garantees locking on right sideband
Loop gain gives:
  Diagnostic of system health
  Accurate loop BW design

## Slide 40: User Interface: Main window
RF Spectrum
Time domain display
IQ diagram
Phase noise Spectrum
Loop filter transfer function
Access to VNA
Raw ADC data recording
Dither lock-in results

## Slide 41: User interface: Frequency counters & output monitoring
PLL Channel 1
Frequency counter
PLL Channel 2
Output monitoring

## Slide 42: Processing Latency
407 ns total latency
Max closed-loop BW:
  1/(8T) ~ 300 kHz
Demod
Loop filter
Freq counter
ADC
DAC
Lockin
PLL Channel 1
FPGA
50 ns
207 ns
407 ns
150 ns

## Slide 43: Feature: Vector Network Analyzer
VNA measures exactly the transfer function that matters
  Exact same TF as seen by loop filter
  Works open-loop or closed-loop!
  Challenging to do with analog PLL
Demod
Loop filter
Freq counter
ADC
DAC
VNA
Lockin
PLL Channel 1
FPGA
D
LPF
Laser
Laser
drive

## Slide 44: VNA Example results 1/2
CEO lock transfer function:
  Current modulation of pump laser

## Slide 45: VNA Example results 2/2
CEO lock transfer function
  Closed-loop measurements yields noise rejection transfer function
~140 dB dyn range limit

## Slide 46: Feature: Frequency counter
Demodulator outputs frequency vs time...
Rectangular or Triangular type*
  : sinc(fT) response
  : sinc2(fT) response
T=1 sec nominal gate time
Zero dead-time
*Ref: Dawkins et al.: Measurement of Stability of Oscillators with Frequency Counters, IEEE Trans. Ultrasonics, Ferroelectrics and freq. control, vol 54, no 5, 2007
DDC
arctan
diff
counter
VCO
to PC
f0+fm
ADC
DAC
Transfer function validation experiment:

## Slide 47: High-level block diagram (1 channel only)
Demod
Loop filter
Freq counter
ADC
DAC
VNA
Lockin
PLL Channel 1
FPGA